## Sheet1
deviceName | deviceHandle | categories | description | cardText | keyFeatures | useCases | connections | softwareConfig | softwareLink | compatibleWith | other | github | bitbucket | bonsaiPackages
Behavior | behaviorDevice | Multipurpose | This is a multi-purpose device tailored to behavioral experiments. It allows control of pokes, RGB LEDs, LEDs, cameras, servo motors and a quadrature counter. | At the heart of a behavioural setup, this device is used to configure and track peripherals such as pokes, LEDs, motors, cameras etc. | - Enable pulse and configure duration for each output - Support for PWM configuration (frequency up to 10KHz and duty cycle) - LED current (up to 100mA) and maximum nominal current configuration - Up to 2 cameras controlling (DO0 and DO1) - Up to 2 servo motors (DO2 and DO3) - Up to 1 quadrature counter (Port2) (Note: Port0 and Port1 can also act as quadrature counters but this is a request feature) | - Behavioral experiments | - 1x clock sync input (CLKIN) [stereo jack] - 1x USB (for computer) [USB type B] - 1x 12V supply [barrel connector jack] - 3x poke connectors (LED+12V valve) (Poke0, Poke1, Poke2) [RJ45] - 2x LEDs outputs (5V) (LED0, LED1) [screw terminal] - 2x RGB LEDs outputs (5V) (RGB0, RGB1) [flick lock 3x male pins] - 4x general purpose digital outputs with PWM (5V) (DO0-DO3) [screw terminal] - 1x ADC (5V) [screw terminal] | - Outputs: pulse duration, PWM (frequency and duty cycle) - LEDs: RGB values, current, and maximum nominal current - Extended features: cameras, servo motors, and quadrature counter - Event firing on: Digital DIs from Ports / Infrared beams from Pokes, Digital DIOs from Ports, Data (includes Analog Input and Encoder), When send a trigger to camera on DO 0, When send a trigger to camera on DO 1 | https://www.google.com/url?q=https%3A%2F%2Fbitbucket.org%2Ffchampalimaud%2Fdownloads%2Fdownloads%2FHarp%2520Behavior%2520v2.0.0.zip&sa=D&sntz=1&usg=AOvVaw1HGd7wVYA0V_2_2wDvRWE7 |  |  | https://github.com/harp-tech/harp_behavior | https://bitbucket.org/fchampalimaud/device.behavior/src/master/
WEAR Basestation | wearBasestation | Sensor Interface | This device is the basestation for the Wired and Wireless sensor devices. The devices can be configured using the Harp Wear software. | This device is the basestation for the Wired and Wireless sensor devices. The devices can be configured using the Harp Wear software. | - Sample rate up to 500 Hz - LED optical stimulation - Start & Stop through pin or software - Output pulse for valves - Accelerometer, Gyroscope, and Magnetometer - Wired or Wireless communication | - Use this device to configure the settings of the WEAR wired or wireless sensor devices (light motion sensors with electrostimulation capabilities) | - 1x clock sync input (CLKIN) [stereo jack] - 1x clock sync output(CLKOUT) [stereo jack] - 1x USB (for computer) [USB type B] - 1x 5V supply [barrel connector jack] - 2x Digital inputs DI0 [BNC], DI1 [screw terminal] - 2x Digital outputs: DO0 [BNC], DO1 [screw terminal] - 1x Analog input: AI0 [screw terminal] - 2x Cameras triggering: CAM0 [flick lock], CAM1 [screw terminal] - 2x Servo motor controllers: CAM0 [flick lock], CAM1 [screw terminal] | - Sensors: sample frequency, accelerometer range, gyroscope range - Sensor device wireless settings: device ID, IR LEDs ON period, TX power, and prioritization - Stimulation: trigger mode (SW-only or input signal dependent), ON/OFF periods, number of cycle repetitions, and current - I/O: input trigger mode and output behavior - Cameras: sample frequency - Servo motors: PWM and pulse periods - Event firing on: Acquisition status, Data arrives from sensor, Any of the inputs is read, Camera is triggered, DO0 each second and pulse ending on output DO1, Sensor metadata, HW and SW Versions of both sensor and receiver | https://www.google.com/url?q=https%3A%2F%2Fbitbucket.org%2Ffchampalimaud%2Fdownloads%2Fdownloads%2FHarp%2520Wear%2520v1.3.4.zip&sa=D&sntz=1&usg=AOvVaw1AFYY-Pa3m_w9ZNoSqKxZa
Camera trigger control | cameraTriggerControl | Digital Interface | This device allows to trigger up to 2 cameras at a predefined frequency pulse or 2 servo motors. Connect cameras to CAM0 TRIG and CAM1 TRIG ports. | Allows up to 2 cameras to be triggered at a predefined frequency | - Separate sample frequency for each camera - Alternate between camera triggering and servo motor controlling - Each camera output can be handled as pure digital output | - Experiments that require camera recordings with external control | - 1x clock sync input (CLKIN) [stereo jack] - 1x USB (for computer) [USB type B] - 1x 5V supply out [screw terminal] - 4x digital outputs (I/O interface for 2 cameras): CAM0 TRIG, CAM0 SYNC, CAM1 TRIG, CAM1 SYNC [screw terminal] - 1x digital input for general purpose or camera control (IN0) [BNC] | - Cameras: sample frequency (up to 1KHz) - Servo motor controller: PWM period (microseconds) and pulse period (microseconds) - Digital input: General purpose or enable camera/motor depending on IN0 state - Sync pulses (up to 100s) - Event firing on: trigger and synchronism outputs, and IN0 | https://www.google.com/url?q=https%3A%2F%2Fbitbucket.org%2Ffchampalimaud%2Fdownloads%2Fdownloads%2FHarp%2520Camera%2520Controller%2520v1.0.1.zip&sa=D&sntz=1&usg=AOvVaw37Pdw8O64-V-p5cilpQj_C |  |  | https://github.com/harp-tech/harp_camera_controller | https://bitbucket.org/fchampalimaud/device.cameracontroller/src/master/
RFID Reader | RFIDreader | Sensor Interface | [9 empty] | https://github.com/harp-tech/harp_RFID_reader
Wear Wireless Motion Sensor | wearWirelessSensor | Sensor Interface | This wireless device has a 9 axis motion sensor sending data in real time to the Wear basestation. It also alllows to add IR Leds for tracking and an LED for optostimulation | Wireless low weight 9-axis motion sensor | - Weight: ~2 g (version 2.1), ~0.8 g (version 3.1) - Configurable sample rate up to 200 Hz (when connected to the Wear Basestation) - Stimulation current up to 50 mA (when connected to the Wear Basestation) - Configurable ranges (when connected to the Wear Basestation) | Appplications where motion needs to be recorded requiring low weight and no wires. | [5 empty] | https://bitbucket.org/fchampalimaud/device.wear.wirelesssensor/src/master/
Load Cells Reader | loadCellsReader | Sensor Interface | This device is an ADC interface for reading load cells values and output to the Load Cells interface device. | Analog interface for 4 load cells sensors | - 1kHz sampling rate - Load cells offset compensation - Up to 4 load cell sensors | - Pressure sensitive experiments that use load cells sensors | - 4x flick lock connectors (X1 to X4) with 4 analog inputs each (5V, I+, I-, 0V) - 1x digital output [RJ45]
Load Cells Interface | loadCellsInterface | Sensor Interface | This device receives digital inputs from 1 or 2 load cells devices and handles data to be output to a computer through a USB port. | Board for interfacing with up two load cells boards | - Controls up to 2 load cells devices - Digital outputs can be regulated at 3V or 5V | - Pressure sensitive experiments that use load cells sensors | - 1x BNC input signal (IN0) - 1x BNC output signal (Out0) - 1x screw terminal connector interface (GND, DOUT1 to DOUT8, DIN0) - 2x RJ45 for digital inputs (from load cells devices) - 1x stereo jack for clock sync input (CLKIN) - 1x USB (for computer) - 1x power barrel connector jack (12V only) - 1x screw terminal connector for 5V output (GND, +5V) | - | https://bitbucket.org/fchampalimaud/downloads/downloads/Harp%20Load%20Cells%20v1.1.0.zip |  |  | https://bitbucket.org/fchampalimaud/device.loadcells/src/master/
LED Array Interface | LEDArrayInterface | Sensor Interface | Controlling LED arrays that require more power. Ideally for optostimulation experiments. | LED array controller for optostimulation experiments | - Controls up to 2 separate RGB LED - Arrays and one common Infrared LED 2 Behavior Mode choices for each RGB LED Array: PWM or Interval | - Opto stimulation experiments that required arrays or matrix of LEDS | - 1x clock sync input (CLKIN) [stereo jack] - 1x USB (for computer) [USB type B] - 1x flat cable - 2x Digital Inputs (IN0, IN1) [flick lock male pins] - 2x Digital Outputs (OUT0, OUT1) [flick lock male pins] - 1x connector for 2 LED Arrays [10-pin flat-cable] | - LED Array PWM: frequency up to 2 KHz, duty cycle, number of pulses - LED Array Interval mode: period ON/OFF, number of pulses, period before repeat, number of repetitions - Digital input - Digital output - LED Array Intensity - Infrared LED intensity - Event firing on: switches toggles, and input (IN0 or IN1) toggles - Control bits: enable power into circuits, start and stop behavior, enable switches, digital outputs, auxiliary digital outputs, and status of the digital inputs | https://bitbucket.org/fchampalimaud/downloads/downloads/Harp%20Load%20Cells%20v1.1.0.zip |  |  | https://bitbucket.org/fchampalimaud/device.ledarray/src/master/
RGB Controller | RGBController | Sensor Interface | Controls 2 arrays of colored LEDs according to RGB scale (Up to 32 LED on each array). Each LED has an associated index. | Strips of RGB LEDs (up tp 32) can be controlled with this board | - Supports two LED array stripes - Each stripe can have up to 32 LEDs - Configurable color and intensity of each LED | Experiments requiring configurable arrays of LEDs, both in intensity and color. | - 1x clock sync input (CLKIN) [stereo jack] - 1x USB (for computer) [USB type C] - 1x 12V supply [barrel connector jack] - 1x I/O interface (5x general purpose digital outputs DO0-DO4, 1x general purpose digital input DI0) [screw terminal] - 1x digital output (DO0) [BNC] - 1x digital input (DI0) [BNC] - 1x Array 0 (G, S, 5V) [flick lock and screw terminal] - 1x Array 1 (G, S, 5V) [flick lock and screw terminal] | - Digital input: general purpose , update LED array with a rising edge, or LED array is ON when pin is high - Digital outputs: controlled by software, 1ms pulse when LED array is updated, 1ms pulse when LED array have a new configuration, toggle when LED array is updated, or toggle when LED array have a new configuration - Event firing on: LED array status (enabled/updated and disabled), and digital input | https://bitbucket.org/fchampalimaud/downloads/downloads/Harp%20RGB%20Array%20v1.0.0.zip |  |  | https://bitbucket.org/fchampalimaud/device.rgb/src/master/
Syringe Pump | syringePump | Sensor Interface
Audio Switch | audioSwitch | Sound | This is a multiplexer device that allows an analog input signal to be forwarded to several output channels. It allows to forward an audio signal to a single speaker or a combination of speakers. The configuration of speakers can be predefined by software or by using a set of digital inputs. | Configurable multiple audio switch | Configuration of up to 15 speakers (depending on the input signal strength) Several speakers can be activated concurrently | Setups that require several speakers to be controlled. | 1x stereo jack for clock sync input (CLKIN) 1x USB (for computer) 1x power barrel connector jack (5V only) 1x screw terminal connector for output power (GND, +5V) 1x screw terminal connector for speaker configuration (GND, DOUT0, GND, DIN0 to DIN4) | - | https://bitbucket.org/fchampalimaud/downloads/downloads/Harp%20Audio%20Switch%20v1.0.0.zip |  |  |  | https://bitbucket.org/fchampalimaud/device.audioswitch/src/master/
Sound Board | soundBoard | Sound | The sound card is a high performance device with 24bit resolution and up to 192kHz sampling rate. The on board flash memory allows a very low latency for the sounds to be played. | High performance sound card | - 24 bit resolution - 2 output channels - Output voltage 2VRMS - 196KHz sampling rate - THD -111dB @ 1kHz 2VRMS - 113 dB SNR (A weighted 80kHz bandwidth) | Auditory tasks that require a low distortion high performance sound systems with low latency | - 3x digital inputs (5V tolerant) - 3x digital outputs (3.3V or 5V) - 2x analog inputs (3.3V max) (5V will not damage the inputs) - 1x serial port (pin compatible with BPod) - 1x digital in/out port (pin compatible with BPod) - 1x micro USB - 1x mini USB - 1x RCA for left channel - 1x RCA for right channel - 1x jack input - 1x 12V supply input - 1x reset button |  | https://bitbucket.org/fchampalimaud/downloads/downloads/Harp%20Sound%20Card%20v1.2.1.zip |  |  | https://bitbucket.org/fchampalimaud/device.soundcard/src/master/
Audio Amplifier | audioAmplifier | Sound | Low distortion, high bandwidth audio amplifier for use in high fidelity applications to drive high quality sounds to speakers | High bandwidth audio amplifier with low distortion and no signal loss over frequency | - 119 dB SNR (A-weighted 80kHz bandwidth) - Unity gain - <0.03% distortion @ 1W - Supports 4Ohms speaker @ 1W - Spectral variation <0.1dB 20Hz - 80kHz | Drive speakers with low distortion and no signal variation over frequency | - 2x 12V supply input - 1x RCA input - 2x 4mm banana outputs - 1x power switch
Analog Input | analogInput | Analog Interface | This harp board support the simultaneous acquisition of up two analog signals with 16 bit resolution and 1kHz sampling | Multiple analog input converter board | - 4 analog input channels - 1kHz sampling frequency - 16 bit resolution - Configurable threshold detection | For any application requiring the acquisition of analog signals | - 4x BNC analog inputs -1x stereo jack for clock sync input (CLKIN) - 1x USB (for computer) - 1x power barrel connector jack - 4x digital outputs (screw terminal) - 1x digital input (screw terminal) - 5V output (screw terminal) |  |  |  |  | https://github.com/harp-tech/harp_analog_input
Multi PWM generator | pwmGenerator | Digital Interface | This board provides 4 independent PWM outputs that can be triggered by four different input triggers or all at the same time. Also four output synchronization signals are available. The PWM signal generation is configured in software. | This board provides 4 independent PWM outputs that can be triggered by four different input triggers or all at the same time. | - PWM output generation - Configurable frequency and duty cycle - Frequency up to 32 KHz - Enabling mechanisms by software or hardware to prevent erroneous triggerings - Start and stop triggers (also by software) - Infinite or configured number of pulses - Complete trigger mechanisms | - Camera triggering - Laser control | - 1x clock sync input (CLKIN) [stereo jack] - 1x USB (for computer) [USB type B] - 1x 12V supply [barrel connector jack] - 5x Digital input triggers: CH0 to CH3, ALL [BNC] - 4x Digital output PWM signals: CH0 to CH3 [BNC] - 5x Digital output sync signals: CH0 to CH3, ALL [BNC] | - PWM outputs: Frequency, duty cycle, number of pulses, Mode (count pulses or infinite) - Triggers: mask and mode (start or start and stop) - Event firing on: start and stop of each PWM burst | https://www.google.com/url?q=https%3A%2F%2Fbitbucket.org%2Ffchampalimaud%2Fdownloads%2Fdownloads%2FHarp%2520Multi%2520Pwm%2520Generator%2520v2.1.0.zip&sa=D&sntz=1&usg=AOvVaw2Wg0b379x9WilnpYdOI8wc |  |  |  | https://bitbucket.org/fchampalimaud/device.multipwm/src/master/
Synchronizer | synchronizer | Digital Interface | Used to synchronize and align the different apparatus and devices present in a complex setup. It timestamps each rising or falling edge of the 9 digital input signals. The sampling of the inputs can also be configured to use a ﬁxed sampling frequency. |  | Receives and timestamps up to 9 digital input signals Sampling on signal transitions or fixed sampling frequency up to 2 KHz | Setups requiring signal alignment | - 1x clock sync input (CLKIN) [stereo jack] - 1x USB (for computer) [USB type B] - 9x digital inputs (DIN0 to DIN5) [BNC], (DIN6 to DIN8) [screw terminal] - 1x digital output (DOUT0) [screw terminal] | - Input read trigger: SW only, any input changes, Rise/Fall edge of DIN0, Frequency of 100, 250, 500, 1000, or 2000 Hz - Output Pin Mode: toggle when a read occur, equal to DIN0, pulses of 1, 2, or 5 ms, pulses of 250, or 500 us | https://bitbucket.org/fchampalimaud/downloads/downloads/Harp%20Synchronizer%20v1.2.0.zip |  |  |  | https://bitbucket.org/fchampalimaud/device.synchronizer/src/master/
Input Expander | inputExpander | Digital Interface | [9 empty] | https://github.com/harp-tech/harp_expander
Output Expander | outputExpander | Digital Interface | [9 empty] | https://github.com/harp-tech/harp_expander
IBL Behavior Control | ibldevice | Multipurpose | [9 empty] | https://github.com/harp-tech/IBL_behavior_control
Clock synchronizer | clockSynchronizer | Timestamp Generators